board info startup

0878
Used the 8 ch power connector. Is WRONG , blow up 1 A fuse Vin+ . replaced.

Big current , more then 1 A on Vdd and Vee.
found that a resistor for all amplifier was wrongly placed . Removed this resistor.
Current now 480 mA ( Vdd and Vee )

Programmed the DAC addresses for the 4 low DAC's .
Temperature can be read
all DAC's can be set
Bias control works.

Programmed the DAC addresses for the 4 upperr DAC's .

All DAC's are responding .

Added shotky diodes over ouput capacitors

Killed I2C level translator

Inverted offset -> patch to apply

enable -3V via the +5V , via 1K wire resistor R43

place R23 10 K

Added R45 to have some load at input

Opto coupler

see spreadsheet

fig 1: Vbias (68V) disabling 5V

fig 2: Vbias , enabling 5 V