The Solid Power distribution considerations

in case this is read from the pdf document check if the date in the footer of the page correspons with the original file date on bit bucket
Original file location : readout system -> cabling -> powersupplies
# General consideration

## description
Assumption here is a 50 plane detector.
Each plane will be equipped with two analogue amplifier boards and two ADC-FPGA boards.
So total 100 analogue and 100 ADC-FPGA boards.
The detector will be built up in sub modules. Here it is assumed that a sub module is 10 planes.

One can distinguish 3 power supply needs:
- For the SiPm bias , this will need a power supply with a voltage between 25 and 85 V this is distributed via the analogue board
- The digital board requires a 5V or 3.3V power supply with reasonable current requirement
- The analogue board need a ~ 7V power supply with low current requirements.

It is not considered to have for the analogue and ADC-FPGA board the same power supply.
## Reliability
The experiment has to run minimal 5 years. max 8years. So reliability calculation have to made for 8 years. The experiment is expected to run 24/ 24 7/7/
( one year == 364 * 24 = 8740 hour )

The power supplies are supposed to be placed that they can be easily accessed.
So initial there could be a trade off between price and reliability ( so take in account to exchange the power supplies at half the life time by example).

Other consideration is the spare policy :
 Do we need a spare for each power supply ( analogue, bias , digital ) ?
- Consider the price in respect to wait for ~ week of delivery in case of a failure.
- Consider segmentation. Price is roughly Euro / W. In case on one power supply one has to invest the full price. In case the power supply is split ( each sub module has its own power supply) one can have one spare with a lower price. To be consider / type ( bias, analogue , digital).

Also it is assumed that the power supplies are efficient enough so they don't need water or forced air cooling.

## general requirements

1. The power supplies have to be compliant with the Belgium "Europese laagspanningsrichtlijn " .
   This is compatible with the EU "The Low Voltage Directive"
   (http://ec.europa.eu/growth/sectors/electrical-engineering/lvd-directive/index_en.htm)
  1. special attention is needed for the kind of mains isollation has to be provided in respect to the grounding requirements ( current to deal with) of the experiment.
2. The total maximum power should not exceed xx W as this is the maximum power that can provided by the current foreseen mains network of Br2.
3. Remote control ? Is it a requirement that the power supply can be switech off remotely.

# Bias power supply
The current per board does mainly depends on the bias control chip ADL5317.
The bias quiescent current depends on the output current. This is assumed to be low << 1 uA / SiPm. So total ~ 10uA for a board. 3mA per control chip is here used.
For t he moment the Hamamatsu S12572-050 is the base line .
This has a break down voltage range from 55 .. 75 V. It is not expected the over-voltage to be supplied will be higher then 3V. So a max voltage of 78V has to be provided to the SiPm's.
The bias controller can not supply voltages above 75V . So 80V is the maximum voltage to supply to the controller chip

So total a 3A @ 80V power supply ( == 240 W ) .

Consideration could be to remove the controller chip on the board :
- Will lower the current requirement from 3A to 100 mA.
- the controller act also as noise reducer but this can be done with capacitors and resistors as the SiPm current is very low but not constant so one has to monitor the bias voltage on teh board.
- the controller act as extra handle to cope with the high spread in the breakdown voltage (according the data sheet ). The later one could be mitigated by have the controler chip on a small PCB for the planes where SiPm's are with Vbr in the low ranges.

# Digital power supply
## general remarks
The currents that the eventual power supplies can apply are relative high so the individual boards should be equipped with a fuse to avoid fire in case a LDO (or DC/DC ) will fail.
## 5 V power distribution base line
As base line the digital boards need a 5V power supply with a current requirement of 6A.
So total 600 A == 3000W .
Distribution of such a current to many boards makes a busbar approach obvious.
But a 600A system has also a lot of disadvantages :
- the 600 A is only a total. So for the last sub module you need only 120 A. So a lot of copper not efficient used.
- One has to consider full shorts . This will give very high peak currents. A peak current of 6000 A will give a force of 144 N between the bars.
- Not easy to find a 5V 600A power supply
- 3000W x .2 ( 80% efficiency) = 600 W , can this be air cooled?

So one could consider to have a star distribution at the 5V power supply to 5 x 120 A and have bus bars of 120A on the sub modules. Still a peak current of 6000A ( defined by the power supply ) is possible.
Alternative use 5 power supplies of each 120A ( TDK-LAMBDA 100A ~ 170Euro RWS-600B-5) . This increase the number of available power supplies and limit the implication with very high currents. But it introduce a careful thinking about the "common" connection in case the power supplies are not close to each other and also ground loops between sub modules.

## 48 V power distribution
### The DC/ DC converter advantage and disadvantages

There are two power consumers :
- The ADC's
- The FPGA

The FPGA has its own 5 V to lower voltage DC /DC converters so this has to run on 5V .
Powering the ADC's with an DC/DC converter will save energy but will increase the cost ( of DC/DC converters ) therefore this is not cost effective option.
### alternative power supply distributions
#### mains
One could use one power supply per plane inside the electronic box limit the distribution current to a minimum. On a first thought not very attractive as
- one has to provide on all metal structures a grounding that agrees with the max mains current (expect to be at least 25A)
  - If the power supplies are outside one could argue that SoLID has only SELV ( Safety Extra Low Voltage) . But the bias > 60V DC and so it is not anymore SELV . The current for the bias is very low so grounding could be still "minimum"
- the own dissipation of the power supply goes into the volume so ~ 15% of 3000W = 450W extra to cool

# Analogue power supply
### +5 V

The voltage on the analogue board will be 4.7V This is done by a LDO. So the power supply voltage before the LDO has to be ~ 5.2 V + 20 % - 10 %.
The +5V input will be the same as for the digital power. A extra board will take care of the filtering of the digital noise.
For now the requirement is ~ 11 A / frame → 110 A / module. So ~ 600W power supply
If we can keep the voltage drop < .05 V a 5 V power supply can do the job as this can provide 5.8V ,
If the voltage drop will be higher most likely the nominal power supply voltage 7.5V , harder to find and in addition the power rating increase to 750W. Cost increas expeted > 2 times

### -3V
Current / board will be ~ 400mA
So total current will be 40 A or 4 A per module.
LDO needs at least -3.3 V on the input. So power supply should be ~ -3.5 V

Alternative
DC/DC from 5 V to 3.3 V on the power cleaning board :
A 5 to 3.3 V DC /DC converter 850mA cost ~ 13 Euro => 50 frames => 650 Euro
Advantage
- simplified cabling cost estmate saving ~ 200 euro ( not including man power installation )
- save the power supply but the 5V power supply will be more expensive.
One could leave out the LDO on the two connected boards ( ~ 6 Euro * 50 ) =300Euro. But a LDO has a better voltage stability. In addition the common mode settings are done with help of the -3V so it has to be stable.

Other disadvantages ( beside the price)
- risk of switching noise
- increase of power as the DC/DC converters are not very efficient ( < 70%)